## Slide 1: Electron Beam Lithography
Université de Technologie de Troyes - avril 2024
Electron Beam Lithography – Franck CARCENAC
LAAS – CNRS 7, avenue du Colonel Roche 31031 TOULOUSE - FRANCE

## Slide 2: Purpose
Overview of E-Beam Lithography (EBL):
  Knowledge, concepts, tools, tips, tricks
  2h … I made choices !
Focus on 1 kind of equipment:
  # UTT one 
Prolegomena

## Slide 3: Outline of presentation
  Introduction
  EBL: basic principals
    Exposure
    Developing
    System composition
    Calibration
    Design recommendations
  EBL: concrete practice
    Driving parameters
    Limitations
    Recommendations
    Sum-up
  EBL nowadays
    Competition at industry level
    Tendencies and job evolution

## Slide 4: Outline of presentation
  Introduction
  EBL: basic principals
    Exposure
    Developing
    System composition
    Calibration
    Design recommendations
  EBL: concrete practice
    Driving parameters
    Limitations
    Recommendations
    Sum-up
  EBL nowadays
    Competition at industry level
    Tendencies and job evolution

## Slide 5: General considerations
Lithography:
  Greek: lithos # stone & graphia # writing
  Shaping/modifying locally a material called a "resist"
  Registering a targeted design as a ghost image inside a material called generally "resist" which will be next developed
  …
Must be though with all the other steps: technology of micro/nano fabrication
  Materials deposition: conductors, insulators, semi-conductors, topological insulators, bio-materials, …
  Materials etching : dry and wet etching, removal, …
  Characterization: physical, electrical, chemical, biological, …, aging, …
  Take always into account:
    The limits of each step
    Start by the last one ... up to integration / packaging !
Introduction

## Slide 6: Lithography … why electron beam ?
  Decrease diffraction # decrease wavelength
  λelectrons at 100keV = 0.0037nm
  Simple to produce, simple to shape
  Maskless (generally)
  But …
  Under vacuum
  Max. resolution few nanometers at 200kV (on bulk substrate)
  Why a so big difference between λ and ultimate resolution ?
  Electrostatic repulsion
  Lenses defects and aberrations
  Electrosensitive materials and their chemistry
Introduction

## Slide 7: Different EBL techniques
Beam type: Gaussian / Variable or fixed shape
  Focused beam (point per point exposure): like a SEM - Gaussian beam
  Non focused beam (area exposure):
    Projection of a variable shape
    Projection of a fixed shape (cell projection): needs a mask with holes
Scan type: Vector / Raster
  Vector scan: only necessary parts are scanned
  Raster scan: all the exposure area is scanned
  Beam is blanked when necessary
Exposure type:
  Fixed stage during exposure … field by field with stitching
  Moving stage during exposure: continuous exposure (e.g.: RAITH FBMS mode - avoid stitching in optical guides)
Introduction

## Slide 8: Different EBL techniques
  Shaped beam - Vector scan - fixed / moving stage
  Cell projection (need a mask)
  Multi-beams, micro-columns
    For mask production
    High resolution (~30nm), high speed … high cost
Direct writing of atomic ideograms realized and observed by AFM.
These 2 ideograms (Yuán Zĭ) mean “atom” in Chinese and are made of iron atoms on a copper surface (D.M. Eigler and Co. in IBM research center of Almaden).
  STM
    Atomic resolution
    Very slow
    Low temperature (mK)
Introduction
  Gaussian beam – vector scan - fixed / moving stage
    Very high resolution … ~8nm
    Intermediate speed & cost

## Slide 9: Outline of presentation
  Introduction
  EBL: basic principals
    Exposure
    Developing
    System composition
    Calibration
    Design recommendations
  EBL: concrete practice
    Driving parameters
    Limitations
    Recommendations
    Sum-up
  EBL nowadays
    Competition at industry level
    Tendencies and job evolution

## Slide 10: Exposure: resists
Electrosensitive material:
  Also sensitive to X-Rays, UV, ions, neutrons, …
  A lot of resists used in UV-lithography are sensitive to electrons
Chemical bonds creation or break:
  Organic resist: 6 – 10eV  secondary electrons
  Inorganic resist: > 50eV  high energy electrons  beam itself  high resolution  dose increase
Positive resist:
  Chemical bonds break   molecular weight (e.g.: PMMA, PMMA/CoMAA, PMGI, ZEP520, CSAR, …)
Negative resist:
  Chemical bonds creation   molecular weight (e.g.: ma-N 2400, PMMA, calixarene, …)
Chemically amplified negative resist:
  Chemical bonds break  acid radical creation
  Thermally activated cross-linking (PEB – Post Exposure Bake)   molecular weight (e.g.: SU8, SAL601, …)
Inorganic materials:
  Phase or crystallinity change, chemical bond break or make a hole: HSQ, WO₃, SiO₂, AlF₃…)
EBL: basic principals

## Slide 11: Exposure: where does it take place ?
http://www.gel.usherbrooke.ca/casino/
EBL: basic principals
Close to electrons’ trajectories

## Slide 12: Exposure: electrons' trajectories
Modification of the molecular weight
Scattering leads to creation of a molecular weight gradient
Additive effects (organic resist)
    proximity effect
EBL: basic principals

## Slide 13: Exposure: Beam Point Spread Function
defects density contribution of backscattering :
Low density #  molecular weight
(for a positive organic resist)
defects density contribution of forward scattering :
High density #    molecular weight
(for a positive organic resist)
Ibackscattered
Ibeam
distance from beam impact
intensity
EBL: basic principals
 2 gaussian curves model

## Slide 14: Outline of presentation
  Introduction
  EBL: basic principals
    Exposure
    Developing
    System composition
    Calibration
    Design recommendations
  EBL: concrete practice
    Driving parameters
    Limitations
    Recommendations
    Sum-up
  EBL nowadays
    Competition at industry level
    Tendencies and job evolution

## Slide 15: Developing
EBL results from lithography + developing
Organic resist:
  No dissolution
  Put in solution:
    Solvent cage around the polymer fragments
    Transport phenomenon in the solution:
      Speed is inversely proportional to Molecular Weight
      Can be difficult in small openings i.e. high resolution: <10nm holes in thick positive tone resist
Inorganic resist:
  Chemical reaction with acid or base (HSQ, WO₃, SiO2)
  Self-developing (AlF₃)
EBL: basic principals

## Slide 16: Developing: resist contrast
Development speed = f(molecular weight, concentration, duration, temperature)
Contrast γ # curve slope
High contrast # high selectivity on molecular weight
  Vertical walls
  Higher process latitude
  Higher resolution
  Less sensitive to proximity effects
Low contrast:
  2.5D lithography / grey scale lithography - e.g. optical microlens fabrication
Log dose
0
1
Log dose
0
1
normalized thickness
normalized thickness
Positive
Negative
EBL: basic principals
Dₙ
Dₚ

## Slide 17: Developing: resist profile
Region of the lowest molecular weight
(positive resist)
Molecular weight increases
(positive resist)
EBL: basic principals

## Slide 18: Developing: proximity effects
Molecular weight modification is additive on organic resists
development threshold Dn or Dp
EBL: basic principals

## Slide 19: Developing: proximity effects
Positive resist
EBL: basic principals
Substrate : Si

## Slide 20: Developing: proximity effects
Molecular weight modification is additive on organic polymers
… and can lead to proximity effect:  resolution
development threshold Dn or Dp
EBL: basic principals

## Slide 21: Developing: proximity effects
Substrate : Si
Positive resist
EBL: basic principals

## Slide 22: Exposure: special …
  2 exposures: 20keV – 2keV
  Negative resist
  Design: 2 superimposed perpendicular gratings
  1 development
  Suspended structures
Tilted view
1st grating exposed at 20keV
2nd grating exposed at 2keV
~100nm
EBL: basic principals

## Slide 23: Outline of presentation
  Introduction
  EBL: basic principals
    Exposure
    Developing
    System composition
    Calibration
    Design recommendations
  EBL: concrete practice
    Driving parameters
    Limitations
    Recommendations
    Sum-up
  EBL nowadays
    Competition at industry level
    Tendencies and job evolution

## Slide 24: System composition
Computer Assisted Design (CAD) software
Electron optics and detectors:
  SEM-like
  SEM imaging capabilities (SE & BSE)
Beam blanker
Pattern generator: for beam scanning control (electronics)
Software: link between user / CAD / beam scanning / beam blanking, …
Minimum configuration
Able to work on one field only
EBL: basic principals

## Slide 25: System composition
Pattern bigger than one field:
  Pattern is cut into fields
  Fields stitching (≤50nm precision on Raith e-line machine)
Need of a LASER interferometer:
  Position precision measurement: 2nm or less (0.15nm)
  Reference for moving
  Allow field calibration: size, rotation & shift
EBL: basic principals

## Slide 26: System composition: synoptic
scanning system
STAGE
sample
scanning data
electron beam
scanning
amplifier
A’: reached position
(mechanical precision  ~1µm)
piezos amplifier
interferometer
mirror
A: position to reach
(precision ~1nm)
A

A’
Stepper motors
piezos. elts
offset calculation
EBL: basic principals

## Slide 27: System composition: remarks
Specimen not always in the axis of mechanical displacement: authorized misalignment 2°
  Alignment registration on sample edge/flat/notch
  Software correction (mechanical precision)
SEM fields are limited in size and position:
  Big field # big distortions
  Not always in the axis of mechanical displacement (electromagnet lenses)
  Be careful: electronic rotation induces distortions (non-linear electronics)
Field size is not calibrated
  Need a mark to calibrate the machine: first level exposure
  Need reference marks made by a calibrated machine: next levels exposure
Electron optics with low distortions
Stepper motors (precision: 1 or few µm)
EBL: basic principals

## Slide 28: Outline of presentation
  Introduction
  EBL: basic principals
    Exposure
    Developing
    System composition
    Calibration
    Design recommendations
  EBL: concrete practice
    Driving parameters
    Limitations
    Recommendations
    Sum-up
  EBL nowadays
    Competition at industry level
    Tendencies and job evolution

## Slide 29: Field calibration
Field is a trapezoid  calibration
  Coordinate system transformation
  Define XY unit vectors value (field size)
  Define angle between XY unit vectors
  Define rotation (allowed 2° … if more image rotation is added)
  Define shift between imaging and lithography scanning ?
3 unknown parameters (field size, rotation, shift): 3 XY mismatch measurements
Can be done by means of:
  Calibrated sample with marks (without LASER interferometry): sample do not move
  Using LASER interferometry: reference point is moved within the field area
EBL: basic principals

## Slide 30: Field calibration: stitching
Before field calibration
After field calibration
With interferometric measurement only
EBL: basic principals

## Slide 31: Outline of presentation
  Introduction
  EBL: basic principals
    Exposure
    Developing
    System composition
    Calibration
    Design recommendations
  EBL: concrete practice
    Driving parameters
    Limitations
    Recommendations
    Sum-up
  EBL nowadays
    Competition at industry level
    Tendencies and job evolution

## Slide 32: Design recommendations: file type - softwares
Drawings’ file type:
  World-famous GDSII (binary files) … DXF, ASCII, …, conversion possible
  DATATYPE parameter used for dose coding (be careful DATATYPE=1000 means DoseFactor=1)
Softwares used:
  Best one: E-Beam writer's software
  If not: must be compatible with the machine
    Klayout (include a PYTHON library … useful for scripting GDS2 files)
    CADENCE, LayoutEditor, CleWin (GDS hierarchy not fully compatible with RAITH)
    GenISys Beamer:
      Process simulation
      Data preparation/correction
      Powerful but expensive (100k€)
EBL: concrete practice

## Slide 33: Design recommendations: drawing/design
Drawing’s grid must match the step-size of exposure
  To prevent coordinates from being rounded by the machine without control
Design is exposed with an orthogonal matrix XY of points:
  All structures non-parallel to X or Y axis are divided in small parts in order to fit the drawing : time and memory consuming
  Try to follow the axes
Avoid nanostructures with stitching
  Better in the center of one field
EBL: concrete practice

## Slide 34: Design recommendations: drawing without knowledge
WHAT IS THE DIFFERENCE BETWEEN THE 2 IMAGES
CLUE :
they have the same scale (µm)
DRAW CLOSE TO THE CENTER
EBL: concrete practice

## Slide 35: Design recommendations: drawing without knowledge
CLUE :
This is not an effect from the display
WHAT IS CURIOUS ON THIS DRAWING ?
Promote X-Y axis parallelism for drawing
If not, this will lead to subdivision of drawing
and exposure time increase.
EBL: concrete practice

## Slide 36: Design recommendations: drawing without knowledge
If not necessary: fit all values to microns or multiple
USE ADAPTED DESIGN GRID(S)
The smallest design grid step should be the exposure step-size:
ASK FOR THE VALUE !
SNAP TO GRID !
SNAP TO ELEMENTS !
EBL: concrete practice

## Slide 37: Design recommendations: drawing without knowledge
WHAT IS THERE BEHIND THE DRAWING ?
50nm
ALWAYS HAVE A LOOK TO DRAWING
WITHOUT THE “FILL” OPTION !
EBL: concrete practice

## Slide 38: Design recommendations: rounding effect
EXPOSURE RESULT CAN BE DIFFERENT FROM THE DRAWING
EBL: concrete practice

## Slide 39: Design recommendations: rounding effect
Circle: Ø 50nm
Exposure step-size: 10nm
Circle: Ø 50nm
Exposure step-size: 2nm
Ø = 20 nm
EBL: concrete practice

## Slide 40: Outline of presentation
  Introduction
  EBL: basic principals
    Exposure
    Developing
    System composition
    Calibration
    Design recommendations
  EBL: concrete practice
    Driving parameters
    Limitations
    Recommendations
    Sum-up
  EBL nowadays
    Competition at industry level
    Tendencies and job evolution

## Slide 41: Driving parameters: design exposure
Py
Px
developed size  beam size x 3 (min.)
beam size
drawing size  desired size – developed beam size
EBL: concrete practice

## Slide 42: Driving parameters: exposure dose
Area dose:
Line dose:
Point dose:
X size
Y size
total number of points
EBL: concrete practice

## Slide 43: Driving parameters: exposure dose
Depends on:
  Substrate
  Energy
  Resist: nature, thickness
  Developer: nature/concentration, duration, mechanical agitation, temperature
  Design: size/pitch/distance between structures (proximity effects), way of exposure, …
Also depends on user’s experience
EBL: concrete practice

## Slide 44: Outline of presentation
  Introduction
  EBL: basic principals
    Exposure
    Developing
    System composition
    Calibration
    Design recommendations
  EBL: concrete practice
    Driving parameters
    Limitations
    Recommendations
    Sum-up
  EBL nowadays
    Competition at industry level
    Tendencies and job evolution

## Slide 45: Limitations: beam diameter = f(current)
EBL: concrete practice
Beam diameter almost constant at "low current" 

## Slide 46: Limitations
Beam size does not depend so much on current …. using a S-FEG gun
"tpoint" limited by scanning generator maximum frequency
"px" define the Line Edge Roughness (LER)
For equipment based on standard scan coils:
  Beam speed: px/tpoint limited by scan coils frequency bandwidth
Line Width Roughness (LWR) and LER depend on noises:
  Mechanical & acoustic vibrations
  Electric/magnetic fields
  Electronics noises,
  Temperature variation …
EBL: concrete practice

## Slide 47: Limitations: beam speed effect
Franck CARCENAC - University of Toulouse - CNRS-LAAS – France franck.carcenac@laas.fr
Origin : scan-coils (inductance)
Induce phase-shift between command and beam position
EBL: concrete practice
Drawing
Drawing
Beam position
Command
Result

## Slide 48: Limitations: beam speed effect
1μm stitching error between 2 polygons induced by a too high beam speed when using SEM scan coils for patterning
Cut the polygons and overlap parts if needed
Distortion at beginning of each line induced by a too high beam speed when using SEM scan coils for patterning
EBL: concrete practice

## Slide 49: Exposure parameters: limiting factors
  Surface too large: exposure time ≈ ∞
    Minimize the surface – moderation of request
    « Mix & Match » - huge surfaces made by photolithography
  Pitch between points: define the Line Edge Roughness (LER)
    Pitch ≈ 1/3 minimum point diameter after developing
    Linked to beam diameter (and developing): if Φbeam  p can 
    Does not change the exposure time: if p  tpoint  (i.e.: F )
    Define the maximum current according to the maximum frequency
  Highest current: to decrease exposure time
    I   reach max. beam speed (~5mm/s)   pitch  max LER  Imax
  Higher sensitive polymer: to decrease exposure time
    Reach Fmax   p if possible or  I   time exposure (interesting ?)
  Decrease the HV (dose=f(HV)): to decrease exposure time
    HV  (beam size , penetration depth )  dose   F  and/or p  (interesting ?)
For equipment based on standard scanning coils
EBL: concrete practice

## Slide 50: Exposure parameters: compromises
Current / beam energy / pitch between points / field size / resist / technology / reasonable time … need experience
Stage movement between fields, beam stabilization time, data transfer, …
  Can double exposure time
Always too long …
  e.g.: 1cm2 of 150nm square fully random with a precision of 10nm on each structure
    10Go data, 100h exposure … on a machine equivalent to UTT one (20keV, 10MHz, 135pA) Because of the combination of resolution and precision, current must be low  exposure time is high.
    1h on a more powerful machine (100keV, 5nA, 125MHz) If scanning generator maximum frequency is higher, current can be increased while maintaining a good resolution
  Collaboration with RENATECH (CNRS public facilities/clean-rooms network) could be a solution for exceptional needs.
EBL: concrete practice

## Slide 51: Extrinsic limitations
Beam drift:
  Room air-conditioning stability: ~ 0.1 - 0.5°C/h
  E-beam writer temperature stability (chiller temperature control: ±0.1°C)
  Mechanical drift: holder/machine temperature (mechanical expansion: ≈ 1µm/0.1°C of variation)
Magnetic fields: powering cables, motors,…
Earth magnetic field: perturbation by moving heavy metal parts (vehicles, lift)
Mechanical vibrations: building, lift, vehicles, pumps, compressor, …, sound
Flat / structured substrate: depth of focus, variable resist thickness, …
EBL: concrete practice

## Slide 52: Intrinsic limitations
Electronics quality:
  Speed
  Noise
  Stability
  Reliability
DAC:
  Number of bits
  Maximum frequency
EBL: concrete practice
Scanning:
  Beam scanning: frequency bandwidth
  Field or field + subfields
  Matrix correction: 4 points or more
  Dead times:
    Stabilization: stage, beam
    Data transfer and treatment
Computer:
  PC under Windows
  Workstation under LINUX
Beam blanking speed
…

## Slide 53: Outline of presentation
  Introduction
  EBL: basic principals
    Exposure
    Developing
    System composition
    Calibration
    Design recommendations
  EBL: concrete practice
    Driving parameters
    Limitations
    Recommendations
    Sum-up
  EBL nowadays
    Competition at industry level
    Tendencies and job evolution

## Slide 54: Recommendations: EBL from A to Z
  Project / design to realize
  Adjustments
  Exposure / development
  Characterization / measurements
  Is it simple ?
Like a SEM column !
Easy to do !
yes … with experience
Do you know how to use the E-beam machine ?
Bibliography, description, bottlenecks, issues …
watch out for proximity effects, total surface, …
Do I know how to use dedicated machines !?!
EBL: concrete practice

## Slide 55: Recommendations: technology integration
Elements to take into account:
Sample :
  Conductor / insulator ?
  Sensitive to electrons ?
  Size / shape / fragility / surface roughness / flatness / available quantity for tests / price / …
Particular properties for the resist and/or sample specificity:
  Previous and next technological steps?
  Sensitivity or resistance towards treatments: wet/dry etching, solubility, compatibility with metals, … ?
  Positive or negative ?
  Final structure: bump or hole ? surface / density / size ?
  Etching or lift-off ?
  Resolution, thickness, dose, developer ?
  Easy to remove ?
Alignment marks : overlay, « mix & match »
EBL: concrete practice

## Slide 56: Outline of presentation
  Introduction
  EBL: basic principals
    Exposure
    Developing
    System composition
    Calibration
    Design recommendations
  EBL: concrete practice
    Driving parameters
    Limitations
    Recommendations
    Sum-up
  EBL nowadays
    Competition at industry level
    Tendencies and job evolution

## Slide 57: Sum-up of SEM - EBL courses
SEM:
  Alignment and astigmatism routines:
    Round shape beam (no astigmatism)
    Minimum size (highest resolution) for the parameters used
  More and more simple:
    Automatic alignment
    Automatic adjustments
  Needs experience to choose the right parameters:
    Share informations
    Try new parameters set: your ability in alignment and astigmatism routines allow you to explore quickly different parameters
  Useful for EBL adjustments
EBL
  Need of experience # need to practice & share knowledges
  Adjustments, calibrations, exposure, developing, characterization … it takes time to be confident
  When facing issues you will learn why and how to overcome them
EBL: concrete practice

## Slide 58: Sum-up manufacturers
Raith
Jeol
ELIONIX
Vistec
Advantest
Crestec
Nanobeam
Nabity: Nano Pattern Generation System
Hitachi
EBL: concrete practice

## Slide 59: Outline of presentation
  Introduction
  EBL: basic principals
    Exposure
    Developing
    System composition
    Calibration
    Design recommendations
  EBL: concrete practice
    Driving parameters
    Limitations
    Recommendations
    Sum-up
  EBL nowadays
    Competition at industry level
    Tendencies and job evolution

## Slide 60: Competition at industry level: 20 years ago …
ELECTRONS : SHAPED BEAM
RESOLUTION - m
MASS PRODUCTION
 NEED MASK
0,001
0,01
0,1
1
10
10-6
10-5
10-4
10-3
10-2
10-1
1
10
102
103
writing speed cm2/s
DIRECT WRITING
MASK-LESS
DUV
ELECTRONS: CELL PROJECTION
X- RAYS ?
EUV
MINIMUM SPEED REQUIRED
FOR PRODUCTION
STM - AFM
ELECTRONS : FOCUSED BEAM
IONS : FOCUSED BEAM
increase speed
increase resolution
?
Nano-Imprint ?
Multiple e-beam ?
EBL nowadays

## Slide 61: International project
EBL nowadays

## Slide 62: Competition at industry level: nowadays !
ELECTRONS : SHAPED BEAM
RESOLUTION - m
MASS PRODUCTION
 NEED MASK
0,001
0,01
0,1
1
10
10-6
10-5
10-4
10-3
10-2
10-1
1
10
102
103
writing speed cm2/s
DIRECT WRITING
MASK-LESS
DUV
ELECTRONS: CELL PROJECTION
X- RAYS ?
MINIMUM SPEED REQUIRED
FOR PRODUCTION
STM - AFM
ELECTRONS : FOCUSED BEAM
IONS : FOCUSED BEAM
increase speed
increase resolution
?
Nano-Imprint ?
Multiple e-beam ?
EUV
EBL nowadays

## Slide 63: Outline of presentation
  Introduction
  EBL: basic principals
    Exposure
    Developing
    System composition
    Calibration
    Design recommendations
  EBL: concrete practice
    Driving parameters
    Limitations
    Recommendations
    Sum-up
  EBL nowadays
    Competition at industry level
    Tendencies and job evolution

## Slide 64: Tendencies: EBL still alive
Still necessary for:
  Optical mask fabrication: UV, DUV, EUV
  Molds fabrication: Nano Imprint Lithography (molding)
Large surfaces:
  Optical metasurfaces: ~1cm2 and more … if possible
    Structures are not regular
    Large surface
    High resolution
    Critical Dimension (CD) control: ~10nm
  Improve machines: current, speed, resolution
  Improve resists: sensitivity, contrast, etching resistance, …
EBL nowadays

## Slide 65: Job evolution
Machines more and more automatic:
  Focus
  Astigmatism
  Calibrations
User: closer and closer to technology
  Scientific project, materials compatibility, flow chart, characterization
  There is still plenty of knowledges to acquire
Scripting (big tendency):
  For complex and variable drawings
  To drive the machine
  Learn programming languages:
    PYTHON, RUBY, MATLAB, …
    Probably deep learning to extract parameters from a collection of exposures
EBL nowadays

## Slide 66: EBL to do what ? Answer some questions …
You search for a lithography solution
R&D / production ?
Large areas exposure - mask making / big samples ?
High throughput or high number of samples / high speed ?
High resolution (nanostructures) ?
Small area : one field exposure (simplest system) ?
Large area : field cannot be as large as the area ?
    Stitching between fields
    High precision stage: moving and control (feedback)
Multi-level alignment : mark detection, overlay ?

## Slide 67: EBL to do what ? Answer some questions …
You want to use an existing machine:
What is your project ?
Did you make some bibliography ?
What are the limits of the machine ? Do they match your requirements ?
Accessibility: for use on your own (needs training), availability
Have you been trained for technology/fabrication ?
Do you need high/low resolution ? Both ?
Do you need precision ? Which one ?
Do you need overlay ? If yes: what about alignment marks ?
Detailed description of your sample ?
Detailed description of previous/following steps ?
Do you need a specific compatibility ?

## Slide 68: Exposure parameters: other limiting factors
  Field stitching can increase dramatically exposure time:
    Mechanical displacements:
    From almost nothing to few seconds: depending on the distance
    Stabilization time: 3s
    Focus correction (automatic): 1s
    Alignment: ~5s
  Proximity effect correction:
    Increase the total settling time by number of objects increase
    To minimize : apply correction only where needed
"Exposure time" : 15s (without PEC)
3mn 20s (with PEC)

## Slide 69: Strategy to speed up EBL: 2 currents & 2 field-size
Possible solution:
    Small structures made at low current low field-size
    Big structures made at high current and large field-size
Needs:
    Define the required total overlay precision (this will define the way to make them)
    Alignment marks (could be made by optical lithography)
  Advantages: everything is made by EBL … only one machine
  Example : if there is not too much area to do
  8 contact pads of 150x150µm2 & nanostructures in 100x100µm2
    High resolution : 20kV, 135pA, 100µm field-size, step-size 10nm
    Low resolution : 20kV, 1-5nA, 800µm field-size, step-size 100nm
Drawback: beam-speed effect for low resolution step

## Slide 70: Automatic Mark Detection: Linescan
Why only automatic ?
  Manual alignment on images:
    you are obliged to be in front of the machine
  Automatic with images:
    Depend on signal
    Precision depend on image size:
      High magnification # small possibility for misalignment
      Low magnification # small precision
    Depend on mark quality
    Observation = exposure: lift-off destroy all the mark
  Automatic mark detection is reliable and efficient

## Slide 71: Automatic Mark detection

## Slide 72: Automatic Mark detection

## Slide 73: Automatic Mark detection

## Slide 74: Automatic Mark detection

## Slide 75: Automatic Mark detection : real sample
1st mark
2nd mark
3rd mark
X
Y

## Slide 76: Automatic Mark detection: real sample
Big differences between example and real sample
  Why ?
    Current used
    HV used
    Mark thickness: low thickness # low contrast
    Recessed mark : no chemical contrast
    Mark material
    Resist material : difficult to see marks under HSQ (contain Si)
    Resist thickness: low Z # low emission + absorption of SE
      High resolution implies low thickness
      Low thickness implies low mark thickness:
        If not : non-uniformity & voids after coating
    Detector: BSED more suitable for metals
    Modification of marks during previous steps

## Slide 77: Automatic Mark detection : real sample
After particle deposition,
mark detection signal is no more the same

## Slide 78: Automatic Mark detection: real sample
Alignment is reliable if scanning across marks
  Determination of the center
  Center position is independent of previous steps : etching, overexposure, …
  increase averaging zone in case of problem
Alignment using edges is not reliable:
  Edge position depend on previous steps
  Lead always to misalignment
  Useful when:
    there is no other solution
    precision is not required
Alignment precision depend on field-size
  Big field-size # less precision due to linear correction of non-linear distortion
  Limit the field-size

## Slide 79: Automatic Mark detection : real sample
After annealing
Before annealing
Mean value of channel position is the same before and after annealing

## Slide 80: Introducing myself
Franck CARCENAC
Involved in electron beam lithography and nanofabrication:
  Since 1984
  Lithographies (UV, X-Rays, Electrons), lift-off, wet/dry etching, characterization
  Si and III-V compounds
  Nanoelectronics, optics, nanomagnetism, nanobiotechnologies among others
  Adaptation, testing/purchasing procedures, installation:
    30keV SEM
    30 - 50keV E-Beam writer (soon 100keV)
    200keV TEM/STEM
People training on SEM, EBL, nanofabrication:
  France and international

## Slide 81: Bibliography
Best jams are made in old jars … today's physics is based on the sum of all previous discoveries
P. Rai-Choudhury (Ed.) : Handbook of Microlithography,
http://en.wikipedia.org/wiki/Electron_beam_lithography
SPIE Handbook of Microlithography, Micromachining and Microfabrication
Nanotechnology, special number of “Microelectronic Engineering”. Vol. 32, September, 1996.
Nanometer-Scale Science & Technology, special number of “Proceedings of the IEEE”. Vol 85, April 1996.
N. Taniguchi (Ed.) Nanotechnology : Integrated Processing Systems for Ultra-precision and Ultra-fine Products, Oxford - University Press, 1997.
Philip Ball, Made to Measure, Princeton University Press - 1994.
H.C.Hoch et al (Eds.), Nanofabrication and Biosystems - Cambridge University Press, 1996.
Semiconductor Industrial Association, The National Technology Roadmap for Semiconductors,1997.
Fundamentals of microfabrication – Marc Madou - CRC Press, (1997)

## Slide 82: Exposure
Electrons from the beam:
  Exposure at beam impact position
  Forward scattering: enlargement of the beam inside the resist
Backscattered electrons:
  Belong to primary beam: energy >50eV
  Escape mainly from substrate through the electrosensitive material
  Enlargement from beam impact position to a distance proportional to energy (~1µm/10keV)
  Trajectory simulation: CASINO free software from Sherbrooke university - CANADA
  ~20% of primary electrons on Si @ 30keV (30% on GaAs, 50% on gold)
Secondary electrons:
  Produced by interaction of high energy electrons with matter: energy <50eV
  Close to scattered electrons trajectory (~10nm)
  ~80% of electrons involved in exposure (on Si)  main contribution to proximity effects